## Slide 1: Ion Torrent™ – Run Analysis and Introduction to the Torrent Browser
Torrent Suite™ v5.0.2

## Slide 2: Objectives for this part of the Training
Introduction of Torrent Suite™ 5.0 Software
Understanding each step of the process
Be able to monitor a run
Be able to review a run report and validate the run
Be able to export the CSA file, and the report file
Be able to reanalyze a run
Notes: Needs to be adapted per session – Here LAB
How do we do that? And Time frames
We will go into the lab!
X
How does this relate to your expectations – trainer to refer to flipcharts with expectations
Anything we need to or can add?
Needs to be in the update loop when content = description will change

## Slide 3: Ion Workflow Overview
Data Analysis
Clonal
Amplification
Isolate Positive
Ion Sphere™
Particles
Load Chip
and Sequence
Prepare
Library
10
10
DNA
A
Sequencing
D
Template Prep
C
Compatible Library Prep
B
Compatible Data
E
Notes: Ion Torrent Systems, Inc.
3

## Slide 5: Log into Torrent browser
3 main tabs
Plan: Set up your run
Monitor: Check your run in real time
Data: Analyze your run report(s)
Site/Server name (configurable)
Notes: How to add the Site/Server name to the Torrent Browser:
FYI: The name is arbitrary and can be whatever the customer would like
Go to Configure
Go to “Customize Site Name” and enter the desired name
Click “Save Name”

## Slide 6
Choose your documentation search method

## Slide 7: Monitor Tab
Information is populated in real time
Notes: Run stays on Monitor tab for 6 hrs. After that, view run in Data tab.
7

## Slide 8: Data Tab – Table View & List View
Table View is useful for scanning,
searching, or sorting a large number of results

## Slide 9: Data Tab – Table View & List View
Search by run name,
run date, chip type, etc.
Click to see run report
Table of Runs
& Run Reports

## Slide 10: Data Tab – Table View & List View
Click to see run report
Open a report

## Slide 11
Thumbnail report, generated fast giving you an estimation of the metrics before the entire analysis is performed
(1/171 of the chip)
960,000 out of 164,699,136 wells
Full report with final metrics

## Slide 12: How to reanalyze a run?
1
2
3

## Slide 13
Unclick Thumbnail

## Slide 14: Run Report - Overview
Barcoding table
(if applicable)
Unaligned/
Aligned metrics
Plugin summaries
Output files
Navigation bar

## Slide 15: Run Report – Top
Navigate run report
See other run reports
for this run*
* Most recent analysis defaults as representative run report.
User can change representative run report.

## Slide 16: Run Report – Barcoding information (if used)
Total output
AQ20 Bases
Mean AQ20 read length
AQ20 reads
For each Barcode, we show

## Slide 17: Run Report – Summary of Unaligned reads
Information about bases and reads before mapping
Efficiency of chip loading
Percentage of wells that
generated reads that
passed filtering
Average read lengths
Notes: 2/2/2017
17

## Slide 18
Chip Loading Image: Very Good Loading Example
Ion PI™ chip heat map,
indicating percent loading across
the physical surface
Notes: 2/2/2017
18

## Slide 19
Chip Loading Image: Very Good Loading Example
Notch
Sample Loading port
Direction of the reagent flows
Notes: 2/2/2017
19

## Slide 20: Run Report– Key Signal
The Consensus Key 1-Mer graph shows the strength of the signal from the first three one-mer bases of the library key. This graph represents the consensus signal measurement of release of H+ during nucleotide incorporation. The y-axis shows signal strength, measured in Counts, which is an arbitrary but consistent unit of measure. The x-axis shows time as nucleotide Flows over the chip.
Notes: 2/2/2017
20

## Slide 21: Run Report – Chip Well Details
Measure Loading Performance and enrichment efficiency
First Pass Counting & Grouping (Library or Test Fragment) Ion Sphere™ Particles
Notes: Please consult the Analysis Report Description for the most accurate explanation of what each of these values represents
2/2/2017
21

## Slide 22: Ion Sphere™ Particle (ISP) Identification Summary
164 699 136
Live ISPs
125 520 074
Lib ISPs
125 406 651
TF ISPs
113 423
Too weak signal
Empty
Wells with ISPs
130 715 177
(key differentiation)
(TCAG)
(ATCG)
Lib reads
80 568 165
Primer dimer
Low quality
Polyclonal
Mouse over for help
Notes: 22

## Slide 23: Ion Sphere™ Particle (ISP) Identification Enrichment Calculation
Lib ISPs
125,406,651
TF ISPs
113,423
Wells with ISPs
130,715,177
% enrich=
-
)
(
Notes: 23

## Slide 24: Run Report – Read Length Details
Histogram of trimmed read length for all reads in output files
Notes: Just an example of some of the graphical outputs in the report. No need to describe details during this presentation. It will be covered during the Torrent Browser training.
FYI the little shoulder peak is normal and is seen following calibration to faster flow. [Chuck S can give a sentence to explain as I didn’t quite get it]
Note that trimming improves overall read quality and may shorten read lengths (lower histogram).
2/2/2017
24

## Slide 25: Run Report – Reads Aligned to Reference
Shows metrics that can be gathered after mapping to reference

## Slide 26: Run report links
*File format of older software version (Torrent Suite™ Software 2.2 or earlier)

## Slide 27: Run Report – Links
File formats produced for extensive data analysis via third party software packages (offered on Webstore) or other bioinformatics tools.
Notes: Show customers how easy it is to access the processed data – you just click on the links! The Excel export function is very useful as well for tracking multiple run metrics over time.
Make sure to point out Support Archive.
2/2/2017
27

## Slide 28: Plugin Summary
variantCaller Features
Supports Ion TargetSeq™, Ion AmpliSeq™, & whole genome applications Each may be run as Germ Line or Somatic
Enables calling for all custom products Supports any genome or custom BED file
Summarizes statistics for barcodes, coverage, chromosomes, variants & hotspots
Calls at 5% or 50%
Compatible with barcoded samples
Planning page for automatic variant calling

## Slide 29: Test Fragment Summary
Notes: 29

## Slide 30: Report Information
Notes: 30

## Slide 31: Report Information
Notes: 31

## Slide 32: Customer Support Archive (CSA)
Files for Troubleshooting;
useful for Technical Support

## Slide 33: Reviewing Data - Projects
Project = set of runs

## Slide 34: About Link
Notes: 34

## Slide 35: References Link
Notes: 35

## Slide 36: Services Link
Notes: 36

## Slide 37: Services Link
Notes: 37

## Slide 38: Configure Link
Notes: 38

## Slide 40: Server and Instrument upgrade
… With Admin rights….. (“Ionadmin” by default)….

## Slide 41: Server and Instrument upgrade
1
2
Options/update menu

## Slide 42: Ion Torrent™ : Semiconductor Sequencing for Life™
Support
Robert David
6178723352
Robert.David@thermofisher
ionsupport@lifetech.com
Call 800-955-6288 in USA
Bioinformatics support
ngs-amsupport@lifetech.com
Jingwei.Ni@thermofisher.com
Online info for Ion Torrent Systems Thermofisher Ion Community
The Chip is the Machine™
Notes: PLEASE CHANGE TECH SUPPORT NUMBER TO MATCH YOUR REGION’S CONTACT INFORMATION – THIS IS US-ONLY

## Slide 43
For Research Use Only. Not for use in diagnostic procedures.
Ion Torrent™ is a Thermo Fisher Scientific brand. © 2014 Thermo Fisher Scientific, Inc. All rights reserved. TaqMan is a registered trademark of Roche Molecular Systems, Inc., used under permission and license. All other trademarks are the property of Thermo Fisher Scientific and its subsidiaries.